# USING A POINTER TIP WITH A 3D ELECTROMAGNETIC DIGITIZER

H.J. Sommer III, Ph.D., Professor of Mechanical Engineering

The Pennsylvania State University

337 Leonhard Building, University Park, PA 16802

(814)863-8997   FAX (814)865-9693

hjs1@psu.edu    http://www.mne.psu.edu/sommer/

# Nomenclature

{xi}	3x1 vector for global position measurement i of digitizer sensor

[Ri]	3x3 rotation matrix for global attitude measurement i of digitizer sensor

{gi}	3x1 vector for global position measurement i of pointer tip using digitizer sensor

{p}	3x1 vector for constant local position of pointer tip with respect to digitizer sensor coordinate frame

{q)	3x1 vector for constant global position of pointer tip within calibration hole

n	number of calibration measurements

[X]	3n x 6 concatenated matrix of rotation measurements [Ri]

{Y}	3n x 1 concatenated vector of position measurements {xi}

{b}	6x1 concatenated vector of {q} and {p}

[I3]	identity matrix of order 3

# Methodolgy

- Rigidly attach the pointer tip to the digitizer sensor.

- Drill a very shallow calibration hole in the top surface of a plastic plate using only the tip of a medium size drill bit.  The pointer tip must be able to touch the bottom of the calibration hole over approximately  +/- 45 degrees of pointer angular motion about the calibration hole.

- Fix the plate rigidly relative to the digitizer transmitter.  The calibration hole should be placed in a region that is representative of future volumes to be digitized

- To calibrate the local pointer tip location, place the pointer tip in the bottom of the calibration hole and rotate the pointer over approximately +/- 45 degrees in all three directions.  Digitize 20 or more calibration samples of {xi} and [Ri] spanning these rotations.  The pointer tip must remain in the bottom of the calibration hole for all calibration samples.

- Calculate the global location of the calibration hole and the local location of the pointer tip with respect to the digitizer sensor using Equation 7.

- Use Equation 1 to predict global locations of the pointer tip for all calibration samples and compare predictions to the global location of the calibration hole to assess measurement precision.

- Use Equation 1 to digitize points with the pointer tip.

# Mathematics

The global pointer tip location {gi} may be calculated using the constant local pointer tip location {p} as shown in Equation 1.  The local location of the pointer tip will be unknown until it is calibrated.

{gi} = {xi} + [Ri]{p}	Eq. 1

During calibration, the global pointer tip location will be approximately equal to the unknown global location of the calibration hole {q} for all samples as shown in Equation 2.  This calibration relationship may be rearranged into the matrix form of Equation 3 and then concatenated for all samples as shown in Equations 4 and 5.

{q}  {gi}          during calibration	Eq. 2

{xi}  {q} – [Ri]{p} =  I3−Ri qp	Eq. 3

x1x2⋮xn≈I3−R1I3−R2⋮⋮I3−Rnqp	Eq. 4

{Y} [X]{b}          for {Y}=x1x2⋮xn     [X]=I3−R1I3−R2⋮⋮I3−Rn     {b}=qp	Eq. 5

The least squares solution for the unknown but constant location of the calibration hole and the unknown but constant local location of the pointer tip are then given in Equations 6 and 7.

{b} = ( [X]T[X] )-1  [X]T{Y}	Eq. 6

qp=nI3−i=1nRi−i=1nRiTnI3−1 i=1nxi−i=1nRiTxi	Eq. 7